SMART ABDUCTION BRACE MONITORING SYSTEM
FOR RELAPSE PREVENTION IN CLUBFOOT

PROBLEM STATEMENT

Clubfoot is a common congenital deformity typically corrected using the Ponseti method. Although initial correction is successful, relapse rates remain high, mainly because of poor compliance with the abduction brace.

Common causes of relapse include:

- Parents not following the prescribed wearing duration

- Incorrect brace fitting (loose straps, improper foot placement)

- Lack of feedback about pressure or discomfort

- No objective monitoring system for clinicians

- Dependence on inaccurate parent-reported usage

Current braces are passive devices with no ability to track:

- Actual wear-time

- Strap tension

- Pressure distribution

- Abduction forces

- Skin comfort conditions

This leads to delayed detection of improper brace use, contributing to preventable relapse.
There is a need for a sensor-based, real-time monitoring system to ensure the brace is worn correctly and consistently.

PRODUCT DESIGN

The proposed solution is an IoT-enabled smart abduction brace monitoring system. Sensors are integrated into a standard Ponseti brace to measure wear-time, pressure distribution, strap tension, temperature and humidity, and movement patterns. This data is transmitted wirelessly to a mobile app and cloud platform for caregiver and clinician monitoring.

SENSOR CONFIGURATION

Pressure sensors (5 sensors)

- Heel pressure sensor: detects heel seating and wear state

- Forefoot pressure sensor: detects forefoot placement and loading

- Midfoot/arch sensor: monitors alignment and fit quality

- Strap tension pressure sensor: identifies loose or overtight straps

- Lateral side pressure sensor: measures abduction forces and foot rotation

Environmental sensor

- Temperature and humidity sensor: monitors the internal microclimate of the brace, helping detect sweating or overheating

Motion sensor

- IMU with accelerometer and gyroscope: detects movement patterns, confirms worn vs not worn conditions, and validates motion data

HARDWARE COMPONENTS

- ESP32 development module with BLE and WiFi
Collects sensor data, processes it, and transmits it wirelessly.

- Five pressure sensors (thin FSR or piezoresistive pads)
Capture pressure distribution across key areas of the brace.

- One IMU (MPU-6050 or similar)
Measures the motion and orientation of the brace.

- One temperature/humidity sensor (SHT31 or DHT22)
Measures internal comfort conditions.

- Analog multiplexer (CD4051) or ADC expander
Expands the number of analog inputs available for the pressure sensors.

- LiPo battery (500–1000 mAh)
Power supply for the wearable unit.

- Battery charging IC (TP4056) with protection
Responsible for the safe charging of the battery.

- Voltage regulator (if required)
Provides stable voltage to sensors and the microcontroller.

- PCB or perf board with connectors
Used to mount and organize components.

- Textile mounting harness or adhesive protective pads
Used to attach sensors inside the brace without discomfort.

- Small enclosure (IP54 rated)
Protects electronics from dust, moisture, and movement.

PROTOTYPE DESCRIPTION

The prototype integrates sensors within the brace and connects them to the ESP32 controller.
Sensors measure pressure, movement, and environmental variables continuously.
The microcontroller preprocesses data, detects correct or incorrect wearing, and transmits results to the parent’s mobile app.
The cloud backend processes long-term data, while the clinician dashboard provides visualizations, summaries, and alerts for relapse prevention.
The system enables caregivers to address issues promptly, enhancing brace compliance and reducing the risk of relapse.